DRAFT Breathing Apparatus Monthly Test - Dräger PSS 100/7000 (v0.1)

NOTE – if you are reading a paper copy of this note it may not be the most up-to-date version. For the latest version view the information on PROCLUS.

Frequency:

- After use (operational or training).

- On the 28th Day after the previous record of monthly inspection

- On acceptance (e.g. returned from servicing/repair, as new equipment, or on return from loan to another HART Unit)

- When replacing a cylinder on a set (e.g. if a cylinder is found to be <280bar on a daily test and needs replacing)

- When attaching a cylinder to a set (this will primarily be during training).

Note: “After use….” This itself must be preceded by a daily test, and can be part of a normal wear during training, or at an incident.

Minimum Competency Level:

- A competent (qualified) BA wearer

- Trainees under the supervision of a Breathing Apparatus Instructor

- Dräger-approved technicians

Personal Hygiene:

- Prior to testing or wearing; all staff must ensure they are clean-shaven around the face seal to achieve an effective fit when using the BA facemask.

- On completion of testing: ensure facemask is appropriately cleaned, using approved cleaning materials as-per the relevant Equipment Data Sheet.

Fault Reporting:

- Any BA set found to be faulty during normal testing must be removed from service immediately. The BA set logbook must be completed to reflect this, including a description of the problem in the fault-reporting section at the back of the logbook. The logbook must then be placed back in its pouch, with the “Do Not Use” sign on the back cover facing out to prevent others inadvertently using the set.

- In all incidences where a BA set fails the HP leak test more than once; the set, facemask, LDV and cylinder must be left in-situ and defected. All parts of the set, cylinder, etc., should remain connected as they were when the HP leak test was carried out, subject to further investigation by a suitably qualified person. Where necessary; the set can be drained of air.

- It must then be reported to the duty Team Leader, as well as your local BA Technician/Champion for them to either repair, or to arrange repair/replacement of the affected items through the manufacturer.

Recording Instructions:

- All records of tests must be recorded in the BA log book immediately on completion of the test, and in BLACK INK. These records must be complete, legible, accurate, and accessible

- Cylinder pressures – when entered – must be rounded-down to the nearest 10bar (e.g., if the gauge reads 309bar; round down to 300bar)

- On completion of a Monthly Test, whether operationally or in-training, an M is to be recorded in the log-book.

Cylinder Inspection

- If attached: remove the cylinder from the BA set

- Examine the cylinder bracket, and connectors on the back-plate, for damage or wear.

- Select cylinder to be attached to set (note: this can be the one removed to check the backplate)

- Ensure cylinder contents ≥ 290bar

- Check cylinder jacket for signs of excessive wear or tear. Ensure, where fitted, Velcro entanglement strap is in good condition, and that the Velcro itself functions correctly.

- Remove the cylinder jacket and inspect cylinder body for damage (any deformity or exposure to carbon layer must be defected).

- Ensure cylinder is within date of first retest by inspecting cylinder label (fig. 1). If cylinder is in use beyond this date – Check cylinder for presence of an in-date re-test sticker similar to that in (Fig. 2).  If in doubt – remove cylinder from use until appropriate documentation can be located.

- Also check that cylinder is within date of final use (fig. 3) – it MUST NOT be used after this date.

(Fig. 1) Date of First Re-Test                                   (Fig. 2) Re-Test Sticker

(Fig. 3) Date of Final Use

- Conduct relevant cleaning.

- Ensure cylinder valve is free from debris

- Refit cylinder jacket

- Connect cylinder, ensuring the connection is finger-tight, and that, where fitted, anti-entanglement straps have also been refitted.

- Now conduct a full inspection and test of the BA set, as per the Daily Test procedure.